| Photometric Test Report |  |  |
| --- | --- | --- |
| Report Number: POTS/GJ13064 | Report Date: 15-02-2013 | Prepared By: G John |
| Test Laboratory: Photometric and Optical Testing Services, Cotswold Business Centre, 42 A P Ellis Road, Rissington Business Park, Upper Rissington, Gloucestershire, GL54 2QB |  |  |
| Company Registration Number: Registered in England & Wales No. OC352911 |  |  |
| Registered Address: Thistle Down Barn, Holcot Lane, Sywell, Northampton, NN6 0BG |  |  |

| Client Details |  |
| --- | --- |
| Company: Astro Lighting | Email: technical@astrolighting.co.uk |
| Address: Astro Lighting Limited, G2 River Way, Harlow CM20 2DP, Great Britain |  |

| Details of Product Tested |  |
| --- | --- |
| Manufacturer: Astro Lighting | Source Type: Domestic light |
| Model: TRIMLESS 12 V LED LAMP | Serial Number: 5623 |
| Lamp Type: LED | Lamp serial number: 1704 |
| Power Supply Used: Uninterruptible AC power supply |  |
| Voltage(AC V) = 243.5 | Current (mA)= 43.5 |
| Power (Watts)= 6.95 | Power factor= 0.656 |

| Integrating Sphere Test |  |  |
| --- | --- | --- |
| Date of Test: 06-02-2013 |  | Ambient Temperature:25°C |
| Measurement Filename: 5624 TRIMLESS 12 V LED LAMP |  |  |
| Instrument Used: Labsphere model CSLMS LED 4060 integrating sphere spectroradiometer |  |  |
| Integrating Sphere Size: 1m |  | Measurement Geometry (2 / 4): 2 |
| Sample Orientation: Horizontal |  | Auxiliary Correction Applied: YES |
| Comments: |  |  |
| Date of Last Calibration (Operating Hours): 08-11-2012 (02:36) |  | Spectral Flux Standard Lamp Used: SCL-1400 |
| Standard Lamp Serial Number: K75 |  | Traceable: to NIST standards |
| Calibration Certificate Number: DM-02008-001 |  | Calibration Certificate Date: 19th February 2010 |
| Calibration Lamp Uncertainty: ± 0.67% (k=2) |  |  |
| Results |  |  |
| Flux (lumens): 272.2 |  | Bare Lamp lumens: 391.8 |
| CIE 1931 Chromaticity Cx: 0.4574 |  | CIE 1931 Chromaticity Cy: 0.4152 |
| CRI (%):84.60 |  | CCT (K): 2768 |
| LOR(%): 69.6 |  |  |

Figure 1: Spectral Irradiance

Figure 2: CIE 1931 diagram.

| Goniophotometer Test |  |  |  |
| --- | --- | --- | --- |
| Date of Test: 06-02-2013 |  | Ambient Temperature: 25°C |  |
| Measurement Filename: 5623 TRIMLESS 12 V LED LAMP |  |  |  |
| Instrument Used: Radiant Imaging NFMS0800 Goniometer with ProMetric PM-1200N-1 Imaging Photometer |  |  |  |
| Photometer Working Distance: 1.5 m |  | Measurement Geometry: Near-Field |  |
| Comments: |  |  |  |
| Reference Photometer Used: Specbos1201 |  | Reference Photometer Serial Number: 2911670 |  |
| Traceable: to NPL standards, UKAS Accredited |  | Calibration Certificate Number: 121104 |  |
| Calibration Certificate Date: 25th October 2011 |  | Sample Stabilisation Time (minutes):75 |  |
| Reference Photometer Calibration Uncertainty: ± 2.4% (k=2, 20-200 lux, CIE illuminant A source) |  |  |  |
| Scan Set Up |  |  |  |
| Direction | Range |  | Increment |
| Inclination Zone 1 | 0-90° |  | 3° |
| Azimuth | 0-360° |  | 10° |
| Results |  |  |  |
| Integrated Luminous Flux (lumens):272.2 | Peak Intensity (3° Spot, candelas): 1010.7 |  | Efficacy (lumens/Watt): 39.45 |
| Beam Angle (50% of max intensity C0-180, degrees): 31.3 |  |  |  |
| Photometric Filename (IES LM-63-2002): 5623 TRIMLESS 12 V LED LAMP |  |  |  |
| IES File – Absolute or Relative Format? RELATIVE |  |  |  |
| Photometric Filename (EULUMDAT): 5623 TRIMLESS 12 V LED LAMP |  |  |  |
| EULUMDAT File – Absolute or Relative Format? RELATIVE |  |  |  |

Figure 3: Far-Field Luminous Intensity (C0-180, Cartesian Coordinates)

Figure 4: Far-Field Luminous Intensity (C0-180, C90-270, Polar Coordinates)

Figure 4. Illuminance cone diagram.

Table 1. UGR values

|  | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 |
| 3 | 976.5 | 978.1 | 975.5 | 972.4 | 974.7 | 975.6 | 974.7 | 978.7 | 977.8 | 981.9 | 985.0 | 988.2 | 989.1 | 993.4 | 997.7 | 999.8 | 1003.2 | 1006.0 | 1007.4 |
| 6 | 916.2 | 919.0 | 911.2 | 913.3 | 915.4 | 918.9 | 922.6 | 925.9 | 933.0 | 937.5 | 944.9 | 948.5 | 954.5 | 960.9 | 968.4 | 974.6 | 980.6 | 988.1 | 990.6 |
| 9 | 790.7 | 795.0 | 794.2 | 796.1 | 808.4 | 816.9 | 823.7 | 828.1 | 844.0 | 854.8 | 867.1 | 875.5 | 890.8 | 905.7 | 916.5 | 926.0 | 936.6 | 947.4 | 952.1 |
| 12 | 582.0 | 576.0 | 579.4 | 599.1 | 596.3 | 613.7 | 647.9 | 657.9 | 679.5 | 702.3 | 722.2 | 745.0 | 775.4 | 788.4 | 811.9 | 840.3 | 852.8 | 869.5 | 883.0 |
| 15 | 345.9 | 342.4 | 339.9 | 369.3 | 360.6 | 376.8 | 420.2 | 424.5 | 455.7 | 484.9 | 514.6 | 537.6 | 592.9 | 606.9 | 640.0 | 690.9 | 704.8 | 734.2 | 753.6 |
| 18 | 170.3 | 167.1 | 175.0 | 171.4 | 182.0 | 192.6 | 205.7 | 234.7 | 249.2 | 276.9 | 299.9 | 333.3 | 348.0 | 386.2 | 420.4 | 452.9 | 501.7 | 521.4 | 545.7 |
| 21 | 78.5 | 76.5 | 79.3 | 82.3 | 85.3 | 91.5 | 100.9 | 111.4 | 122.0 | 139.6 | 151.5 | 168.1 | 185.3 | 203.1 | 228.0 | 255.0 | 278.5 | 294.8 | 314.6 |
| 24 | 37.7 | 36.4 | 37.4 | 40.2 | 43.2 | 46.2 | 49.6 | 54.6 | 58.0 | 67.2 | 72.2 | 81.9 | 90.6 | 102.5 | 115.2 | 125.0 | 131.2 | 137.2 | 147.8 |
| 27 | 19.5 | 19.1 | 20.1 | 20.9 | 21.5 | 22.9 | 25.2 | 27.4 | 28.3 | 32.0 | 34.9 | 39.5 | 43.2 | 46.3 | 52.0 | 58.5 | 62.3 | 61.9 | 64.9 |
| 30 | 11.5 | 11.5 | 11.4 | 12.3 | 12.4 | 13.1 | 14.4 | 14.6 | 15.4 | 16.4 | 17.9 | 18.5 | 21.0 | 21.9 | 24.2 | 28.0 | 28.3 | 30.1 | 30.8 |
| 33 | 7.7 | 7.6 | 7.7 | 7.7 | 8.1 | 8.6 | 8.9 | 9.4 | 9.5 | 9.9 | 10.3 | 10.6 | 10.7 | 11.7 | 12.6 | 13.6 | 15.2 | 16.0 | 16.7 |
| 36 | 5.5 | 5.3 | 5.3 | 5.2 | 5.3 | 5.8 | 6.2 | 6.5 | 6.5 | 6.9 | 7.2 | 7.5 | 7.4 | 7.6 | 7.9 | 8.6 | 9.4 | 9.8 | 10.4 |
| 39 | 3.6 | 3.5 | 3.2 | 3.2 | 3.2 | 3.6 | 4.0 | 4.1 | 4.4 | 4.8 | 5.4 | 5.5 | 5.5 | 5.5 | 5.8 | 6.0 | 6.2 | 6.5 | 6.7 |
| 42 | 2.1 | 2.1 | 1.8 | 1.8 | 1.6 | 2.0 | 2.3 | 2.4 | 2.7 | 3.1 | 3.6 | 3.7 | 3.7 | 3.7 | 3.9 | 4.0 | 4.1 | 4.0 | 4.1 |
| 45 | 0.7 | 1.2 | 0.3 | 0.5 | 0.6 | 0.6 | 0.7 | 1.4 | 1.5 | 1.8 | 2.1 | 2.2 | 2.2 | 2.2 | 2.4 | 2.4 | 2.4 | 2.3 | 2.3 |
| 48 | 0.3 | 0.0 | 0.0 | 0.0 | 0.1 | 0.2 | 0.0 | 0.3 | 0.7 | 0.9 | 0.9 | 1.2 | 1.2 | 0.7 | 1.3 | 1.2 | 1.4 | 1.0 | 1.4 |
| 51 | 0.0 | 0.0 | 0.2 | 0.0 | 0.0 | 0.2 | 0.0 | 0.3 | 0.6 | 0.0 | 0.2 | 0.3 | 0.4 | 0.3 | 0.8 | 0.8 | 0.5 | 0.6 | 0.0 |
| 54 | 0.0 | 0.0 | 0.0 | 0.0 | 0.1 | 0.0 | 0.0 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 0.0 | 0.1 | 0.3 | 0.0 | 0.0 | 0.2 | 0.1 |
| 57 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.1 | 0.0 | 0.0 | 0.2 | 0.0 |
| 60 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.2 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 63 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 66 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 69 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 72 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 75 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 78 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 81 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 84 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 87 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 90 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2a. Luminous intensity values, azimuth 0-180°

|  | 190 | 200 | 210 | 220 | 12 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 | 1003.2 |
| 3 | 1009.6 | 1010.4 | 1010.6 | 1010.7 | 1009.8 | 1008.2 | 1007.1 | 1005.4 | 1001.9 | 997.5 | 996.6 | 992.4 | 990.5 | 986.2 | 981.7 | 981.7 | 981.2 |
| 6 | 995.3 | 994.9 | 995.9 | 995.2 | 993.1 | 990.1 | 986.3 | 983.0 | 975.8 | 969.3 | 960.7 | 954.4 | 946.9 | 938.3 | 931.0 | 922.9 | 924.8 |
| 9 | 959.1 | 961.1 | 960.3 | 959.6 | 956.2 | 948.7 | 938.7 | 931.1 | 918.1 | 904.0 | 884.9 | 875.0 | 861.6 | 846.7 | 826.3 | 815.3 | 804.8 |
| 12 | 889.5 | 891.3 | 893.4 | 881.4 | 872.4 | 868.3 | 847.5 | 827.4 | 803.6 | 773.1 | 742.1 | 723.1 | 681.5 | 653.4 | 643.4 | 610.3 | 591.5 |
| 15 | 762.8 | 758.1 | 766.3 | 736.9 | 721.2 | 719.2 | 673.0 | 647.8 | 607.2 | 561.1 | 513.8 | 497.2 | 441.0 | 414.0 | 412.1 | 367.3 | 356.7 |
| 18 | 551.6 | 559.3 | 532.8 | 518.6 | 497.2 | 467.4 | 455.3 | 409.7 | 369.0 | 320.7 | 291.3 | 247.7 | 226.4 | 209.2 | 191.9 | 190.0 | 174.2 |
| 21 | 315.1 | 318.2 | 310.9 | 292.8 | 271.6 | 256.1 | 235.5 | 205.9 | 180.8 | 149.9 | 131.2 | 115.3 | 102.0 | 93.8 | 89.7 | 84.6 | 79.2 |
| 24 | 145.4 | 147.3 | 150.5 | 145.9 | 132.9 | 118.9 | 106.7 | 89.4 | 80.3 | 67.0 | 61.3 | 54.5 | 50.3 | 46.6 | 43.6 | 39.9 | 37.8 |
| 27 | 64.2 | 67.8 | 67.6 | 63.5 | 58.5 | 55.1 | 48.8 | 41.1 | 37.8 | 33.4 | 31.2 | 28.1 | 25.4 | 23.7 | 23.0 | 21.8 | 19.8 |
| 30 | 31.1 | 30.6 | 32.3 | 30.4 | 28.8 | 28.2 | 24.4 | 22.1 | 20.5 | 19.2 | 17.7 | 16.8 | 15.2 | 14.2 | 13.9 | 12.5 | 11.9 |
| 33 | 16.9 | 16.9 | 16.5 | 17.0 | 16.2 | 15.0 | 14.8 | 13.5 | 13.0 | 12.8 | 12.3 | 11.3 | 10.6 | 9.9 | 9.1 | 8.7 | 7.9 |
| 36 | 10.5 | 10.6 | 10.8 | 11.1 | 10.8 | 10.3 | 10.3 | 9.2 | 9.1 | 9.2 | 9.0 | 8.1 | 7.3 | 6.8 | 6.6 | 6.3 | 5.6 |
| 39 | 6.9 | 6.8 | 7.4 | 7.9 | 8.0 | 7.7 | 7.3 | 6.7 | 6.1 | 6.4 | 5.8 | 5.2 | 4.4 | 4.3 | 4.4 | 4.2 | 3.8 |
| 42 | 4.2 | 4.4 | 4.9 | 5.1 | 5.4 | 5.3 | 4.9 | 4.3 | 3.6 | 3.7 | 3.3 | 3.0 | 2.3 | 2.4 | 2.7 | 2.6 | 2.2 |
| 45 | 2.3 | 2.6 | 2.9 | 3.1 | 3.3 | 3.3 | 2.9 | 2.4 | 1.9 | 1.9 | 1.8 | 1.6 | 1.3 | 1.4 | 1.6 | 1.6 | 1.0 |
| 48 | 1.3 | 1.4 | 1.5 | 1.7 | 1.9 | 1.9 | 1.6 | 1.3 | 0.8 | 0.7 | 0.6 | 0.9 | 0.6 | 0.8 | 1.0 | 0.0 | 0.0 |
| 51 | 0.6 | 0.8 | 0.4 | 0.3 | 0.7 | 0.0 | 0.0 | 0.2 | 0.1 | 0.2 | 0.4 | 0.0 | 0.2 | 0.4 | 0.0 | 0.0 | 0.0 |
| 54 | 0.5 | 0.6 | 0.3 | 0.1 | 0.2 | 0.0 | 0.0 | 0.3 | 0.1 | 0.2 | 0.2 | 0.0 | 0.1 | 0.1 | 0.0 | 0.0 | 0.1 |
| 57 | 0.0 | 0.2 | 0.2 | 0.1 | 0.3 | 0.0 | 0.0 | 0.2 | 0.0 | 0.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 60 | 0.0 | 0.0 | 0.2 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 63 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 66 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 69 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 72 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 75 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 78 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 81 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 84 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 87 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 90 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2b. Luminous intensity values, azimuth 190-350°

Signature:

_________________________________

Print Name:

GH JOHN

_________________________________

Date:

15-02-2013

_________________________________

Partner / Director
Duly authorised to sign on behalf of:

Photometric and Optical Testing Services LLP