| Photometric Test Report |  |  |
| --- | --- | --- |
| Report Number: POTS/GJ13066 | Report Date: 15-02-2013 | Prepared By: G John |
| Test Laboratory: Photometric and Optical Testing Services, Cotswold Business Centre, 42 A P Ellis Road, Rissington Business Park, Upper Rissington, Gloucestershire, GL54 2QB |  |  |
| Company Registration Number: Registered in England & Wales No. OC352911 |  |  |
| Registered Address: Thistle Down Barn, Holcot Lane, Sywell, Northampton, NN6 0BG |  |  |

| Client Details |  |
| --- | --- |
| Company: Astro Lighting | Email: technical@astrolighting.co.uk |
| Address: Astro Lighting Limited, G2 River Way, Harlow CM20 2DP, Great Britain |  |

| Details of Product Tested |  |
| --- | --- |
| Manufacturer: Astro Lighting | Source Type: Domestic light |
| Model: GOYA LED MATT NICKEL | Serial Number: 0873 |
| Lamp Type: LED |  |
| Power Supply Used: Uninterruptible AC power supply |  |
| Voltage(AC V) = 246.0 | Current (mA)= 79 |
| Power (Watts)= 7.10 | Power factor= 0.3652 |

| Integrating Sphere Test |  |  |
| --- | --- | --- |
| Date of Test: 07-02-2013 |  | Ambient Temperature:25°C |
| Measurement Filename: 0873 GOYA LED MATT NICKEL |  |  |
| Instrument Used: Labsphere model CSLMS HALOGEN 4060 integrating sphere spectroradiometer |  |  |
| Integrating Sphere Size: 1m |  | Measurement Geometry (2 / 4): 2 |
| Sample Orientation: Horizontal |  | Auxiliary Correction Applied: YES |
| Comments: |  |  |
| Date of Last Calibration (Operating Hours): 08-11-2012 (02:36) |  | Spectral Flux Standard Lamp Used: SCL-1400 |
| Standard Lamp Serial Number: K75 |  | Traceable: to NIST standards |
| Calibration Certificate Number: DM-02008-001 |  | Calibration Certificate Date: 19th February 2010 |
| Calibration Lamp Uncertainty: ± 0.67% (k=2) |  |  |
| Results |  |  |
| Flux (lumens): 431.1 |  |  |
| CIE 1931 Chromaticity Cx: 0.4239 |  | CIE 1931 Chromaticity Cy: 0.3990 |
| CRI (%):71.1 |  | CCT (K): 3189 |

Figure 1: Spectral Irradiance

Figure 2: CIE 1931 diagram.

| Goniophotometer Test |  |  |  |
| --- | --- | --- | --- |
| Date of Test: 07-02-2013 |  | Ambient Temperature: 25°C |  |
| Measurement Filename: 0873 GOYA LED MATT NICKEL |  |  |  |
| Instrument Used: Radiant Imaging NFMS0800 Goniometer with ProMetric PM-1200N-1 Imaging Photometer |  |  |  |
| Photometer Working Distance: 1.5 m |  | Measurement Geometry: Near-Field |  |
| Comments: |  |  |  |
| Reference Photometer Used: Specbos1201 |  | Reference Photometer Serial Number: 2911670 |  |
| Traceable: to NPL standards, UKAS Accredited |  | Calibration Certificate Number: 121104 |  |
| Calibration Certificate Date: 25th October 2011 |  | Sample Stabilisation Time (minutes):75 |  |
| Reference Photometer Calibration Uncertainty: ± 2.4% (k=2, 20-200 lux, CIE illuminant A source) |  |  |  |
| Scan Set Up |  |  |  |
| Direction | Range |  | Increment |
| Inclination Zone 1 | 0-90° |  | 3° |
| Azimuth | 0-360° |  | 10° |
| Results |  |  |  |
| Integrated Luminous Flux (lumens):431.1 | Peak Intensity (3° Spot, candelas): 169.8 |  | Efficacy (lumens/Watt): 60.72 |
| Beam Angle (50% of max intensity C0-180, degrees): 111.0 |  |  |  |
| Photometric Filename (IES LM-63-2002): 0873 GOYA LED MATT NICKEL |  |  |  |
| IES File – Absolute or Relative Format? RELATIVE |  |  |  |
| Photometric Filename (EULUMDAT): 0873 GOYA LED MATT NICKEL |  |  |  |
| EULUMDAT File – Absolute or Relative Format? RELATIVE |  |  |  |

Figure 3: Far-Field Luminous Intensity (C0-180, Cartesian Coordinates)

Figure 4: Far-Field Luminous Intensity (C0-180, C90-270, Polar Coordinates)

Figure 4. Illuminance cone diagram.

Table 1. UGR values

|  | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 |
| 3 | 166.1 | 166.1 | 166.1 | 165.7 | 166.2 | 167.1 | 167.4 | 165.9 | 165.4 | 163.8 | 163.6 | 163.4 | 164.3 | 165.0 | 166.7 | 168.2 | 169.7 | 169.8 | 169.3 |
| 6 | 164.1 | 162.2 | 163.6 | 166.2 | 165.2 | 163.1 | 164.1 | 164.5 | 164.5 | 166.6 | 166.3 | 164.7 | 162.8 | 163.3 | 166.6 | 165.2 | 162.8 | 164.8 | 165.1 |
| 9 | 161.8 | 160.4 | 162.5 | 162.1 | 160.3 | 162.3 | 159.8 | 160.2 | 161.8 | 160.9 | 159.9 | 162.4 | 159.3 | 159.2 | 159.1 | 160.8 | 159.6 | 159.6 | 161.3 |
| 12 | 157.4 | 161.5 | 160.2 | 155.5 | 159.5 | 157.3 | 156.4 | 158.3 | 157.3 | 155.7 | 160.9 | 156.4 | 157.9 | 158.8 | 159.4 | 160.5 | 159.4 | 162.4 | 163.9 |
| 15 | 158.7 | 161.2 | 155.2 | 156.8 | 155.2 | 154.0 | 155.6 | 155.1 | 154.0 | 154.0 | 154.4 | 156.4 | 154.0 | 154.2 | 155.0 | 157.0 | 158.5 | 156.4 | 162.6 |
| 18 | 154.3 | 155.8 | 151.1 | 154.2 | 152.2 | 150.9 | 150.3 | 151.5 | 147.7 | 151.0 | 150.3 | 152.0 | 152.7 | 151.9 | 157.3 | 154.0 | 152.5 | 151.7 | 159.8 |
| 21 | 150.0 | 154.7 | 149.2 | 149.4 | 147.7 | 150.0 | 147.1 | 145.9 | 149.2 | 148.6 | 147.8 | 146.3 | 145.7 | 145.7 | 150.2 | 147.2 | 153.6 | 147.6 | 158.4 |
| 24 | 149.2 | 149.8 | 147.5 | 146.6 | 145.6 | 141.1 | 144.6 | 141.2 | 142.9 | 143.2 | 141.3 | 142.1 | 141.5 | 144.7 | 145.6 | 141.4 | 148.2 | 147.6 | 148.4 |
| 27 | 144.9 | 146.8 | 141.0 | 139.1 | 140.0 | 137.6 | 134.6 | 138.6 | 138.6 | 135.9 | 138.1 | 138.1 | 136.3 | 137.3 | 140.1 | 139.7 | 145.3 | 137.9 | 146.3 |
| 30 | 141.1 | 139.3 | 134.9 | 132.1 | 136.2 | 134.1 | 133.5 | 130.7 | 131.4 | 131.1 | 131.4 | 130.8 | 133.7 | 133.2 | 133.5 | 138.4 | 136.6 | 139.0 | 144.0 |
| 33 | 135.8 | 134.2 | 127.9 | 129.2 | 126.3 | 129.3 | 127.4 | 124.0 | 126.8 | 125.0 | 126.0 | 126.0 | 130.4 | 126.9 | 129.5 | 132.5 | 132.6 | 130.5 | 138.8 |
| 36 | 128.6 | 126.5 | 123.3 | 122.7 | 124.0 | 119.6 | 119.4 | 121.1 | 121.0 | 118.1 | 120.0 | 119.7 | 120.4 | 120.9 | 119.9 | 128.7 | 123.3 | 125.4 | 134.0 |
| 39 | 124.4 | 120.5 | 116.7 | 120.9 | 112.9 | 113.4 | 112.3 | 115.2 | 113.3 | 112.7 | 113.2 | 114.8 | 116.9 | 114.0 | 116.5 | 118.9 | 114.3 | 119.1 | 129.7 |
| 42 | 116.0 | 113.3 | 114.6 | 107.4 | 109.8 | 109.6 | 107.9 | 108.4 | 108.6 | 104.1 | 107.5 | 107.2 | 105.4 | 111.6 | 112.9 | 111.4 | 111.6 | 114.3 | 119.5 |
| 45 | 110.3 | 104.8 | 104.7 | 106.6 | 102.6 | 100.7 | 100.9 | 98.4 | 99.0 | 98.2 | 100.4 | 98.5 | 101.9 | 106.6 | 102.8 | 103.2 | 101.5 | 104.9 | 112.3 |
| 48 | 106.3 | 97.8 | 95.3 | 95.1 | 97.7 | 95.2 | 91.1 | 92.7 | 90.8 | 92.2 | 90.6 | 93.6 | 93.2 | 95.3 | 95.8 | 95.8 | 95.6 | 100.0 | 104.6 |
| 51 | 94.7 | 89.2 | 91.6 | 88.3 | 86.3 | 84.3 | 82.8 | 86.9 | 83.0 | 82.2 | 84.8 | 87.0 | 85.7 | 85.8 | 88.1 | 88.5 | 86.6 | 90.5 | 97.6 |
| 54 | 82.0 | 83.0 | 84.6 | 81.0 | 78.5 | 78.6 | 74.2 | 76.7 | 76.5 | 73.8 | 76.1 | 76.3 | 76.8 | 78.4 | 79.4 | 78.8 | 78.6 | 84.1 | 89.8 |
| 57 | 80.6 | 72.7 | 73.9 | 72.2 | 71.9 | 68.8 | 69.1 | 68.0 | 68.5 | 65.3 | 69.7 | 68.5 | 70.5 | 69.3 | 68.8 | 68.8 | 72.2 | 75.9 | 79.9 |
| 60 | 70.0 | 67.8 | 64.4 | 63.9 | 59.8 | 58.6 | 61.5 | 57.6 | 58.4 | 56.0 | 59.4 | 59.2 | 60.0 | 61.5 | 61.7 | 62.6 | 64.4 | 69.3 | 72.9 |
| 63 | 63.6 | 58.5 | 54.6 | 53.1 | 49.4 | 51.4 | 49.7 | 48.4 | 47.1 | 41.5 | 42.6 | 49.0 | 49.0 | 52.7 | 53.1 | 52.0 | 57.1 | 60.3 | 63.6 |
| 66 | 52.5 | 51.8 | 47.3 | 46.7 | 43.6 | 40.8 | 35.6 | 31.1 | 26.7 | 21.2 | 24.8 | 29.3 | 33.1 | 41.4 | 43.0 | 45.3 | 49.5 | 52.4 | 55.8 |
| 69 | 44.0 | 40.9 | 41.6 | 39.4 | 33.5 | 25.7 | 21.4 | 15.8 | 7.6 | 5.3 | 6.6 | 9.9 | 17.3 | 26.8 | 29.2 | 36.3 | 40.2 | 45.4 | 48.5 |
| 72 | 37.5 | 36.2 | 31.7 | 28.0 | 20.7 | 8.1 | 5.6 | 3.4 | 1.4 | 0.5 | 1.9 | 1.4 | 6.7 | 8.8 | 16.0 | 28.5 | 32.9 | 36.0 | 39.0 |
| 75 | 29.3 | 25.3 | 21.7 | 21.4 | 8.5 | 1.6 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 5.7 | 16.8 | 25.3 | 27.4 | 29.5 |
| 78 | 19.6 | 17.0 | 16.5 | 6.6 | 1.2 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.3 | 4.5 | 14.7 | 19.5 | 22.0 |
| 81 | 12.6 | 10.2 | 6.2 | 0.7 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 4.8 | 11.4 | 14.1 |
| 84 | 5.9 | 4.2 | 0.4 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 3.6 | 6.8 |
| 87 | 2.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.2 | 0.3 | 2.9 |
| 90 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2a. Luminous intensity values, azimuth 0-180°

|  | 190 | 200 | 210 | 220 | 12 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 | 163.3 |
| 3 | 167.8 | 167.2 | 166.7 | 165.7 | 165.3 | 165.8 | 165.7 | 166.0 | 166.0 | 167.3 | 167.8 | 167.4 | 167.4 | 166.7 | 165.9 | 165.9 | 165.9 |
| 6 | 163.1 | 163.8 | 166.2 | 165.9 | 163.5 | 163.3 | 165.7 | 165.7 | 166.2 | 168.8 | 169.8 | 169.3 | 167.4 | 166.0 | 166.0 | 167.7 | 167.8 |
| 9 | 161.5 | 167.3 | 165.9 | 165.3 | 163.7 | 165.5 | 165.4 | 162.5 | 167.3 | 165.5 | 167.4 | 168.0 | 167.8 | 167.5 | 162.8 | 165.6 | 165.9 |
| 12 | 159.4 | 167.7 | 165.9 | 165.8 | 168.0 | 164.3 | 165.2 | 162.9 | 165.8 | 166.6 | 166.5 | 165.0 | 164.0 | 163.6 | 163.0 | 161.0 | 165.2 |
| 15 | 157.5 | 161.9 | 159.8 | 161.6 | 161.7 | 164.4 | 164.8 | 164.4 | 164.2 | 164.0 | 160.7 | 163.0 | 166.7 | 160.9 | 164.3 | 161.0 | 157.9 |
| 18 | 154.1 | 161.6 | 158.6 | 161.9 | 161.9 | 164.7 | 163.4 | 162.8 | 160.8 | 160.5 | 160.3 | 164.9 | 158.8 | 160.5 | 158.3 | 156.7 | 159.9 |
| 21 | 152.1 | 153.9 | 160.2 | 156.2 | 156.8 | 159.4 | 158.9 | 160.5 | 160.0 | 162.3 | 160.0 | 158.8 | 158.0 | 153.8 | 157.1 | 153.6 | 155.8 |
| 24 | 148.0 | 151.7 | 158.0 | 152.3 | 156.7 | 156.1 | 159.4 | 155.1 | 157.3 | 157.4 | 154.7 | 158.3 | 152.5 | 158.1 | 151.9 | 150.6 | 151.3 |
| 27 | 146.1 | 145.8 | 149.2 | 149.2 | 149.8 | 153.4 | 154.4 | 154.0 | 153.8 | 153.9 | 156.1 | 151.9 | 150.8 | 145.8 | 148.9 | 147.6 | 146.7 |
| 30 | 145.1 | 144.2 | 148.0 | 148.7 | 148.3 | 152.3 | 147.8 | 152.1 | 149.0 | 150.3 | 147.9 | 151.0 | 143.6 | 148.0 | 145.2 | 141.1 | 141.9 |
| 33 | 139.6 | 137.0 | 140.5 | 143.4 | 143.6 | 145.1 | 142.0 | 143.3 | 146.5 | 144.1 | 140.3 | 146.6 | 141.2 | 140.5 | 141.2 | 139.3 | 134.8 |
| 36 | 134.2 | 130.3 | 138.0 | 136.6 | 137.3 | 141.8 | 142.8 | 137.4 | 142.3 | 141.6 | 140.0 | 136.9 | 137.4 | 135.3 | 131.5 | 131.7 | 128.7 |
| 39 | 124.3 | 122.3 | 130.9 | 135.1 | 133.5 | 131.8 | 134.7 | 135.3 | 135.4 | 136.0 | 138.7 | 131.6 | 130.4 | 133.8 | 132.4 | 125.1 | 122.4 |
| 42 | 123.6 | 119.2 | 121.2 | 128.9 | 128.9 | 126.6 | 127.6 | 131.9 | 130.6 | 132.2 | 129.4 | 130.7 | 126.5 | 123.9 | 122.1 | 123.3 | 114.5 |
| 45 | 113.9 | 112.7 | 114.1 | 121.5 | 124.7 | 121.6 | 121.1 | 124.6 | 124.9 | 124.2 | 123.3 | 120.2 | 118.1 | 121.5 | 116.2 | 111.0 | 112.9 |
| 48 | 109.8 | 106.7 | 106.9 | 112.4 | 117.4 | 117.0 | 115.4 | 118.7 | 117.6 | 116.3 | 115.2 | 118.3 | 114.7 | 110.9 | 108.4 | 102.8 | 105.6 |
| 51 | 99.2 | 104.4 | 100.7 | 102.1 | 108.9 | 109.4 | 108.5 | 108.9 | 109.8 | 109.1 | 109.5 | 108.1 | 103.6 | 104.8 | 104.6 | 98.5 | 95.9 |
| 54 | 92.0 | 93.6 | 92.2 | 97.9 | 99.1 | 102.6 | 103.8 | 102.7 | 100.9 | 103.2 | 102.6 | 98.4 | 94.6 | 96.6 | 94.9 | 92.4 | 87.1 |
| 57 | 86.4 | 84.8 | 87.8 | 86.0 | 91.8 | 96.3 | 96.1 | 96.4 | 93.9 | 95.4 | 92.2 | 92.9 | 90.4 | 87.0 | 88.1 | 84.8 | 81.9 |
| 60 | 73.8 | 79.5 | 81.3 | 82.7 | 79.8 | 84.4 | 83.6 | 82.7 | 81.4 | 81.2 | 83.1 | 84.0 | 81.7 | 79.3 | 77.4 | 75.0 | 73.4 |
| 63 | 66.7 | 70.3 | 72.3 | 72.9 | 72.0 | 68.2 | 66.0 | 62.8 | 57.0 | 56.0 | 63.7 | 68.4 | 71.8 | 73.4 | 69.5 | 66.4 | 68.2 |
| 66 | 62.9 | 63.0 | 62.3 | 64.3 | 61.2 | 49.3 | 38.2 | 27.1 | 24.8 | 24.7 | 33.0 | 44.3 | 60.1 | 66.7 | 62.2 | 59.2 | 55.8 |
| 69 | 52.1 | 53.4 | 52.0 | 51.0 | 42.2 | 25.1 | 8.9 | 5.2 | 4.8 | 5.7 | 9.8 | 18.5 | 32.1 | 53.2 | 54.6 | 50.9 | 49.6 |
| 72 | 43.4 | 43.1 | 43.8 | 31.4 | 11.8 | 3.8 | 0.3 | 0.4 | 0.0 | 0.0 | 0.0 | 3.4 | 3.4 | 23.2 | 44.0 | 44.7 | 42.3 |
| 75 | 33.8 | 36.9 | 31.6 | 9.5 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 4.8 | 26.1 | 35.4 | 32.3 |
| 78 | 24.7 | 26.8 | 9.5 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 4.7 | 23.0 | 23.3 |
| 81 | 17.2 | 12.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 5.9 | 16.0 |
| 84 | 8.8 | 2.1 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 6.9 |
| 87 | 3.7 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.5 |
| 90 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2b. Luminous intensity values, azimuth 190-350°

Signature:

_________________________________

Print Name:

GH JOHN

_________________________________

Date:

15-02-2013

_________________________________

Partner / Director
Duly authorised to sign on behalf of:

Photometric and Optical Testing Services LLP