| Photometric Test Report |  |  |
| --- | --- | --- |
| Report Number: POTS/GJ13106 | Report Date: 18-03-2013 | Prepared By: G John |
| Test Laboratory: Photometric and Optical Testing Services, Cotswold Business Centre, 42 A P Ellis Road, Rissington Business Park, Upper Rissington, Gloucestershire, GL54 2QB |  |  |
| Company Registration Number: Registered in England & Wales No. OC352911 |  |  |
| Registered Address: Thistle Down Barn, Holcot Lane, Sywell, Northampton, NN6 0BG |  |  |

| Client Details |  |
| --- | --- |
| Company: Astro Lighting | Email: technical@astrolighting.co.uk |
| Address: Astro Lighting Limited, G2 River Way, Harlow CM20 2DP, Great Britain |  |

| Details of Product Tested |  |
| --- | --- |
| Manufacturer: Astro Lighting | Source Type: Domestic light |
| Model: VETRO TWIN | Serial Number: 5669 |
| Lamp Type: LED |  |
| Power Supply Used: Uninterruptible AC power supply |  |
| Voltage(AC V) = 242.4 | Current (mA)= 114 |
| Power (Watts)= 25.35 | Power factor= 0.9164 |

| Integrating Sphere Test |  |  |
| --- | --- | --- |
| Date of Test: 06-03-2013 |  | Ambient Temperature:25°C |
| Measurement Filename: 5669 VETRO TWIN |  |  |
| Instrument Used: Labsphere model CSLMS HALOGEN 4060 integrating sphere spectroradiometer |  |  |
| Integrating Sphere Size: 1m |  | Measurement Geometry (2 / 4): 4 |
| Sample Orientation: Horizontal |  | Auxiliary Correction Applied: YES |
| Comments: |  |  |
| Date of Last Calibration (Operating Hours): 08-11-2012 (02:36) |  | Spectral Flux Standard Lamp Used: SCL-1400 |
| Standard Lamp Serial Number: K75 |  | Traceable: to NIST standards |
| Calibration Certificate Number: DM-02008-001 |  | Calibration Certificate Date: 19th February 2010 |
| Calibration Lamp Uncertainty: ± 0.67% (k=2) |  |  |
| Results |  |  |
| Flux (lumens): 1191.0 |  |  |
| CIE 1931 Chromaticity Cx: 0.4314 |  | CIE 1931 Chromaticity Cy: 0.3999 |
| CRI (%):82.01 |  | CCT (K): 3059 |

Figure 1: Spectral Irradiance

Figure 2: CIE 1931 diagram.

| Goniophotometer Test |  |  |  |
| --- | --- | --- | --- |
| Date of Test: 06-03-2013 |  | Ambient Temperature: 25°C |  |
| Measurement Filename: 5669 VETRO TWIN |  |  |  |
| Instrument Used: Radiant Imaging NFMS0800 Goniometer with ProMetric PM-1200N-1 Imaging Photometer |  |  |  |
| Photometer Working Distance:1.5 m |  | Measurement Geometry: Near-Field |  |
| Comments: |  |  |  |
| Reference Photometer Used: Specbos1201 |  | Reference Photometer Serial Number: 2911670 |  |
| Traceable: to NPL standards, UKAS Accredited |  | Calibration Certificate Number: 121104 |  |
| Calibration Certificate Date: 25th October 2011 |  | Sample Stabilisation Time (minutes):45 |  |
| Reference Photometer Calibration Uncertainty: ± 2.4% (k=2, 20-200 lux, CIE illuminant A source) |  |  |  |
| Scan Set Up |  |  |  |
| Direction | Range |  | Increment |
| Inclination Zone 1 | 0-90° |  | 3° |
| Azimuth | 0-360° |  | 10° |
| Results |  |  |  |
| Integrated Luminous Flux (lumens):1191 | Peak Intensity (3° Spot, candelas): 4257.0 |  | Efficacy (lumens/Watt): 46.98 |
| Beam Angle (50% of max intensity C0-180, degrees): 20.2 |  |  |  |
| Photometric Filename (IES LM-63-2002): 5669 VETRO TWIN |  |  |  |
| IES File – Absolute or Relative Format? Absolute |  |  |  |
| Photometric Filename (EULUMDAT): 5669 VETRO TWIN |  |  |  |
| EULUMDAT File – Absolute or Relative Format? Absolute |  |  |  |

Figure 3: Far-Field Luminous Intensity (C0-180, Cartesian Coordinates)

Figure 4: Far-Field Luminous Intensity (C0-180, C90-270, Polar Coordinates)

Figure 4. Illuminance cone diagram.

Table 1. UGR values

|  | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 |
| 3 | 4065.9 | 4039.7 | 3997.7 | 4009.9 | 3959.4 | 3927.4 | 3910.5 | 3888.7 | 3866.3 | 3857.2 | 3833.0 | 3821.9 | 3810.2 | 3793.5 | 3796.5 | 3838.8 | 3794.2 | 3833.8 | 3858.8 |
| 6 | 3388.6 | 3334.9 | 3360.3 | 3324.3 | 3259.3 | 3196.3 | 3179.4 | 3132.2 | 3094.8 | 3069.7 | 3039.4 | 3029.8 | 3024.9 | 2995.5 | 3024.8 | 3058.2 | 3064.7 | 3016.7 | 3064.6 |
| 9 | 2543.8 | 2479.4 | 2483.6 | 2393.1 | 2368.1 | 2341.6 | 2315.3 | 2290.5 | 2265.2 | 2231.8 | 2210.3 | 2182.2 | 2156.7 | 2136.9 | 2127.5 | 2118.0 | 2181.3 | 2157.8 | 2211.6 |
| 12 | 1820.5 | 1777.1 | 1746.4 | 1726.1 | 1749.8 | 1706.1 | 1668.3 | 1625.2 | 1599.8 | 1577.4 | 1549.5 | 1527.8 | 1526.8 | 1529.0 | 1542.1 | 1496.3 | 1496.0 | 1510.5 | 1542.7 |
| 15 | 1317.0 | 1284.5 | 1290.7 | 1267.7 | 1248.9 | 1239.0 | 1207.7 | 1190.5 | 1156.8 | 1144.2 | 1118.9 | 1114.1 | 1099.6 | 1103.8 | 1090.8 | 1090.5 | 1097.2 | 1083.0 | 1105.4 |
| 18 | 991.1 | 994.6 | 965.1 | 965.3 | 936.5 | 929.2 | 910.4 | 899.9 | 891.5 | 871.1 | 866.4 | 849.7 | 838.4 | 836.0 | 825.2 | 836.5 | 827.5 | 844.6 | 837.2 |
| 21 | 779.6 | 778.7 | 768.3 | 764.8 | 749.2 | 734.1 | 727.6 | 724.0 | 711.8 | 699.3 | 696.4 | 691.7 | 681.5 | 674.5 | 675.6 | 677.7 | 673.8 | 677.5 | 674.2 |
| 24 | 643.1 | 640.3 | 635.3 | 620.6 | 619.8 | 615.0 | 607.0 | 605.8 | 599.0 | 592.4 | 589.8 | 585.8 | 577.8 | 577.1 | 575.3 | 569.0 | 575.0 | 576.6 | 575.8 |
| 27 | 556.8 | 555.1 | 549.0 | 545.9 | 545.3 | 542.7 | 538.2 | 533.7 | 530.6 | 526.2 | 523.8 | 519.4 | 517.5 | 516.7 | 515.8 | 511.3 | 511.3 | 516.0 | 515.2 |
| 30 | 500.0 | 498.9 | 496.7 | 496.5 | 494.2 | 490.8 | 489.6 | 487.1 | 483.9 | 479.2 | 478.3 | 475.2 | 470.6 | 466.0 | 466.7 | 466.3 | 465.9 | 469.1 | 469.5 |
| 33 | 451.0 | 451.7 | 448.5 | 447.0 | 445.4 | 440.6 | 437.7 | 431.1 | 429.9 | 425.3 | 423.4 | 417.9 | 412.4 | 405.7 | 408.4 | 409.4 | 407.3 | 409.9 | 412.5 |
| 36 | 382.4 | 381.9 | 373.6 | 374.0 | 368.1 | 365.3 | 361.0 | 347.5 | 345.6 | 341.8 | 335.5 | 330.8 | 327.5 | 316.5 | 315.4 | 319.8 | 310.0 | 315.5 | 323.8 |
| 39 | 277.2 | 272.1 | 271.9 | 255.3 | 255.6 | 258.8 | 249.3 | 245.6 | 238.6 | 232.9 | 228.0 | 228.0 | 216.1 | 206.6 | 197.6 | 190.0 | 199.3 | 199.2 | 211.5 |
| 42 | 158.6 | 149.9 | 153.8 | 148.7 | 149.5 | 149.1 | 145.1 | 137.1 | 134.3 | 131.6 | 127.6 | 125.4 | 122.7 | 110.7 | 106.3 | 96.9 | 99.5 | 101.0 | 113.0 |
| 45 | 71.0 | 65.4 | 68.2 | 72.8 | 68.3 | 70.5 | 68.9 | 64.9 | 64.8 | 64.4 | 62.5 | 60.9 | 59.6 | 51.7 | 47.3 | 44.6 | 43.2 | 46.2 | 53.6 |
| 48 | 28.6 | 25.9 | 29.6 | 32.8 | 28.6 | 30.1 | 29.7 | 30.3 | 30.5 | 29.2 | 29.3 | 29.4 | 27.4 | 23.5 | 21.7 | 21.6 | 20.5 | 22.3 | 25.1 |
| 51 | 13.1 | 11.6 | 13.2 | 14.8 | 13.3 | 13.5 | 13.6 | 15.7 | 16.4 | 12.9 | 14.4 | 14.1 | 13.2 | 11.6 | 12.0 | 11.7 | 9.8 | 12.3 | 12.1 |
| 54 | 7.4 | 4.8 | 8.0 | 7.0 | 7.3 | 7.2 | 6.9 | 9.8 | 10.1 | 6.1 | 7.5 | 7.7 | 7.2 | 6.6 | 7.3 | 6.8 | 4.5 | 7.4 | 6.6 |
| 57 | 2.7 | 1.9 | 5.2 | 4.9 | 5.0 | 4.8 | 4.6 | 6.5 | 6.7 | 3.1 | 4.8 | 5.0 | 4.1 | 3.9 | 3.4 | 4.8 | 3.1 | 4.8 | 3.8 |
| 60 | 2.7 | 2.4 | 3.9 | 2.3 | 2.4 | 2.1 | 2.0 | 3.8 | 4.5 | 2.5 | 2.5 | 1.5 | 3.5 | 3.0 | 1.9 | 2.2 | 2.4 | 2.3 | 2.7 |
| 63 | 2.2 | 1.6 | 2.6 | 2.7 | 2.4 | 1.6 | 0.6 | 2.2 | 1.8 | 1.0 | 2.5 | 2.1 | 1.2 | 1.8 | 1.8 | 1.1 | 0.5 | 1.1 | 1.6 |
| 66 | 1.5 | 0.9 | 0.9 | 2.5 | 1.3 | 1.0 | 0.7 | 1.0 | 0.0 | 1.4 | 1.6 | 0.0 | 1.5 | 1.7 | 2.3 | 0.9 | 0.8 | 0.5 | 0.5 |
| 69 | 1.0 | 0.0 | 1.3 | 0.5 | 0.8 | 0.8 | 1.1 | 0.4 | 0.0 | 0.9 | 0.5 | 1.4 | 1.1 | 0.9 | 1.3 | 0.9 | 1.2 | 0.0 | 0.5 |
| 72 | 0.9 | 0.0 | 0.9 | 0.0 | 0.7 | 0.0 | 0.7 | 0.4 | 0.0 | 0.5 | 1.0 | 0.8 | 0.4 | 2.2 | 0.7 | 0.0 | 0.4 | 0.0 | 0.5 |
| 75 | 0.5 | 0.0 | 1.5 | 1.5 | 0.7 | 0.4 | 1.5 | 0.0 | 0.9 | 0.5 | 0.4 | 0.0 | 0.5 | 2.0 | 0.7 | 0.0 | 1.0 | 0.0 | 0.5 |
| 78 | 1.4 | 0.5 | 0.4 | 0.0 | 0.5 | 1.0 | 0.5 | 0.0 | 0.0 | 1.0 | 1.8 | 0.0 | 0.0 | 0.7 | 0.0 | 0.0 | 0.4 | 0.0 | 0.0 |
| 81 | 0.9 | 0.4 | 0.7 | 0.5 | 0.5 | 0.5 | 0.7 | 0.0 | 0.3 | 0.5 | 0.0 | 0.4 | 0.3 | 0.5 | 0.5 | 0.0 | 0.0 | 0.4 | 0.5 |
| 84 | 0.0 | 0.7 | 0.7 | 0.5 | 0.0 | 0.3 | 0.0 | 0.0 | 0.3 | 0.5 | 0.4 | 0.0 | 0.0 | 0.7 | 0.5 | 0.0 | 0.0 | 0.0 | 0.4 |
| 87 | 0.5 | 0.0 | 0.0 | 0.0 | 0.9 | 0.3 | 0.3 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.3 | 0.0 | 0.0 | 0.0 | 0.0 | 0.4 | 0.0 |
| 90 | 0.6 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.4 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2a. Luminous intensity values, azimuth 0-180°

|  | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 | 4257.0 |
| 3 | 3868.7 | 3863.4 | 3944.3 | 3929.7 | 3957.4 | 3995.0 | 4022.5 | 4044.4 | 4070.1 | 4079.6 | 4092.8 | 4100.2 | 4096.9 | 4097.9 | 4119.7 | 4070.2 | 4076.2 |
| 6 | 3068.6 | 3171.3 | 3218.2 | 3227.6 | 3233.6 | 3295.5 | 3326.8 | 3351.5 | 3389.4 | 3412.3 | 3438.8 | 3463.7 | 3451.1 | 3478.9 | 3499.1 | 3477.6 | 3392.4 |
| 9 | 2205.6 | 2282.7 | 2260.1 | 2314.4 | 2359.5 | 2413.3 | 2463.8 | 2506.2 | 2532.7 | 2564.1 | 2579.6 | 2587.0 | 2586.0 | 2579.3 | 2557.8 | 2602.4 | 2536.0 |
| 12 | 1544.5 | 1563.9 | 1599.3 | 1682.4 | 1697.1 | 1719.4 | 1738.0 | 1769.5 | 1799.9 | 1817.6 | 1835.7 | 1869.6 | 1891.5 | 1912.5 | 1850.3 | 1831.5 | 1821.0 |
| 15 | 1101.0 | 1137.6 | 1153.4 | 1175.4 | 1211.3 | 1223.1 | 1253.6 | 1265.7 | 1294.9 | 1303.4 | 1333.1 | 1341.2 | 1362.0 | 1351.3 | 1347.7 | 1345.3 | 1311.3 |
| 18 | 854.5 | 849.3 | 873.4 | 872.3 | 896.1 | 911.0 | 936.9 | 964.5 | 972.5 | 994.3 | 998.0 | 1000.7 | 1008.7 | 1001.0 | 1016.4 | 998.9 | 1012.5 |
| 21 | 682.9 | 687.2 | 700.8 | 705.0 | 708.7 | 725.2 | 747.0 | 758.4 | 765.5 | 781.1 | 793.0 | 790.8 | 788.4 | 795.9 | 801.6 | 792.9 | 791.4 |
| 24 | 581.0 | 584.7 | 583.4 | 594.3 | 599.2 | 605.4 | 619.4 | 627.4 | 633.8 | 642.9 | 650.1 | 647.6 | 651.3 | 652.7 | 646.8 | 654.3 | 650.6 |
| 27 | 519.8 | 518.9 | 523.1 | 529.7 | 532.4 | 537.1 | 540.6 | 547.3 | 551.6 | 556.5 | 559.0 | 561.9 | 564.2 | 566.2 | 563.0 | 561.5 | 564.1 |
| 30 | 472.5 | 474.1 | 477.7 | 479.5 | 479.9 | 485.6 | 490.2 | 494.2 | 494.8 | 499.3 | 502.1 | 500.8 | 499.8 | 504.4 | 505.4 | 502.5 | 504.9 |
| 33 | 416.2 | 420.1 | 423.4 | 426.5 | 424.2 | 427.5 | 431.0 | 437.9 | 437.6 | 440.8 | 441.8 | 442.1 | 442.3 | 449.0 | 450.0 | 446.9 | 451.0 |
| 36 | 327.5 | 331.3 | 341.4 | 344.5 | 342.1 | 345.4 | 344.7 | 352.7 | 354.9 | 355.3 | 358.0 | 364.2 | 365.3 | 369.8 | 374.1 | 366.9 | 375.8 |
| 39 | 213.4 | 224.6 | 218.6 | 234.0 | 234.4 | 233.8 | 240.9 | 244.4 | 245.2 | 247.5 | 257.3 | 256.4 | 261.1 | 259.4 | 256.6 | 269.0 | 270.0 |
| 42 | 112.5 | 120.8 | 122.4 | 138.5 | 133.6 | 136.2 | 134.3 | 138.4 | 139.3 | 140.5 | 149.1 | 156.3 | 152.3 | 156.3 | 155.2 | 161.4 | 155.7 |
| 45 | 53.4 | 57.0 | 61.0 | 67.3 | 67.1 | 68.2 | 65.4 | 68.2 | 67.8 | 68.0 | 74.2 | 80.1 | 72.7 | 77.1 | 82.6 | 79.7 | 72.9 |
| 48 | 26.4 | 28.5 | 29.5 | 30.3 | 33.0 | 32.1 | 31.2 | 33.9 | 31.5 | 32.4 | 35.7 | 37.6 | 32.0 | 37.2 | 41.0 | 36.5 | 31.1 |
| 51 | 14.8 | 14.1 | 14.8 | 15.0 | 17.7 | 15.5 | 15.0 | 19.1 | 15.2 | 17.2 | 17.6 | 18.3 | 15.4 | 20.2 | 18.6 | 15.3 | 15.0 |
| 54 | 9.1 | 8.7 | 7.9 | 8.3 | 10.6 | 8.1 | 8.4 | 11.2 | 8.0 | 9.6 | 10.1 | 9.8 | 8.4 | 11.6 | 8.7 | 7.4 | 8.1 |
| 57 | 5.8 | 4.6 | 5.3 | 5.4 | 6.5 | 4.5 | 5.2 | 6.9 | 5.0 | 6.1 | 6.4 | 6.1 | 5.4 | 7.4 | 4.4 | 2.7 | 4.6 |
| 60 | 2.5 | 1.7 | 4.0 | 2.5 | 3.3 | 1.6 | 3.2 | 3.0 | 2.1 | 2.8 | 3.7 | 3.9 | 4.1 | 4.2 | 2.4 | 1.2 | 2.0 |
| 63 | 1.7 | 1.9 | 2.0 | 1.2 | 1.7 | 1.9 | 2.1 | 1.2 | 0.5 | 1.2 | 2.3 | 1.2 | 2.4 | 3.5 | 2.5 | 0.5 | 1.9 |
| 66 | 1.0 | 1.4 | 1.8 | 0.9 | 0.0 | 1.2 | 2.7 | 0.5 | 0.5 | 1.0 | 1.6 | 0.7 | 1.2 | 1.0 | 1.6 | 1.5 | 0.5 |
| 69 | 1.0 | 1.4 | 1.6 | 1.3 | 0.0 | 1.2 | 0.7 | 0.4 | 1.5 | 0.9 | 1.9 | 0.6 | 1.4 | 0.0 | 0.0 | 0.7 | 1.0 |
| 72 | 0.0 | 0.9 | 0.5 | 1.1 | 0.0 | 0.7 | 0.8 | 0.0 | 0.5 | 0.7 | 0.4 | 0.0 | 0.9 | 1.9 | 1.5 | 1.1 | 0.5 |
| 75 | 0.0 | 0.7 | 0.5 | 0.7 | 0.0 | 1.5 | 0.5 | 0.3 | 0.0 | 0.3 | 0.5 | 0.4 | 1.6 | 0.7 | 0.5 | 1.0 | 0.9 |
| 78 | 1.0 | 0.4 | 0.5 | 0.0 | 0.0 | 1.4 | 0.4 | 0.3 | 0.4 | 0.3 | 0.4 | 0.4 | 1.2 | 2.0 | 1.0 | 0.4 | 0.9 |
| 81 | 0.5 | 0.4 | 0.5 | 0.5 | 0.4 | 0.7 | 0.4 | 0.0 | 0.0 | 0.3 | 0.4 | 0.0 | 0.5 | 0.5 | 0.0 | 0.0 | 0.0 |
| 84 | 0.4 | 0.0 | 0.0 | 0.0 | 0.3 | 0.3 | 0.5 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.5 | 0.5 | 0.0 | 0.4 |
| 87 | 0.4 | 0.5 | 0.0 | 0.0 | 0.0 | 0.4 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |
| 90 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 | 0.0 |

Table 2b. Luminous intensity values, azimuth 190-350°

Signature:

_________________________________

Print Name:

GH JOHN

_________________________________

Date:

___18-03-2013______________________________

Partner / Director
Duly authorised to sign on behalf of:

Photometric and Optical Testing Services LLP