| Photometric Test Report |  |  |
| --- | --- | --- |
| Report Number: POTS/GJ13185 | Report Date: 30-05-2013 | Prepared By: G John |
| Test Laboratory: Photometric and Optical Testing Services, Cotswold Business Centre, 42 A P Ellis Road, Rissington Business Park, Upper Rissington, Gloucestershire, GL54 2QB |  |  |
| Company Registration Number: Registered in England & Wales No. OC352911 |  |  |
| Registered Address: Thistle Down Barn, Holcot Lane, Sywell, Northampton, NN6 0BG |  |  |

| Client Details |  |
| --- | --- |
| Company: Astro Lighting | Email: technical@astrolighting.co.uk |
| Address: Astro Lighting Limited, G2 River Way, Harlow CM20 2DP, Great Britain |  |

| Details of Product Tested |  |
| --- | --- |
| Manufacturer: Astro Lighting | Source Type: Domestic light |
| Model: MASHIKO LED 600 | Serial Number: 7134 |
| Lamp Type: LED |  |
| Power Supply Used: Uninterruptible AC power supply |  |
| Voltage(AC V) = 243.3 | Current (mA)= 98 |
| Power (Watts)= 10.6 | Power factor= 0.446 |

| Integrating Sphere Test |  |  |
| --- | --- | --- |
| Date of Test: 30-05-2013 |  | Ambient Temperature:25°C |
| Measurement Filename: 7134 MASHIKO LED 600 |  |  |
| Instrument Used: Labsphere model CSLMS HALOGEN 4060 integrating sphere spectroradiometer |  |  |
| Integrating Sphere Size: 1m |  | Measurement Geometry (2 / 4): 4 |
| Sample Orientation: Horizontal |  | Auxiliary Correction Applied: YES |
| Comments: |  |  |
| Date of Last Calibration (Operating Hours): 02-05-2013 (02:42) |  | Spectral Flux Standard Lamp Used: SCL-1400 |
| Standard Lamp Serial Number: K75 |  | Traceable: to NIST standards |
| Calibration Certificate Number: DM-02008-001 |  | Calibration Certificate Date: 19th February 2010 |
| Calibration Lamp Uncertainty: ± 0.67% (k=2) |  |  |
| Results |  |  |
| Flux (lumens):657.8 |  |  |
| CIE 1931 Chromaticity Cx: 0.4304 |  | CIE 1931 Chromaticity Cy: 0.3949 |
| CRI (%):88.28 |  | CCT (K): 3034 |

Figure 1: Spectral Irradiance

Figure 2: CIE 1931 diagram.

| Goniophotometer Test |  |  |  |
| --- | --- | --- | --- |
| Date of Test: 29-05-2013 |  | Ambient Temperature: 25°C |  |
| Measurement Filename: 7134 MASHIKO LED 600 |  |  |  |
| Instrument Used: Radiant Imaging NFMS0800 Goniometer with ProMetric PM-1200N-1 Imaging Photometer |  |  |  |
| Photometer Working Distance:3 m |  | Measurement Geometry: Near-Field |  |
| Comments: |  |  |  |
| Reference Photometer Used: Specbos1201 |  | Reference Photometer Serial Number: 2911670 |  |
| Traceable: to NPL standards, UKAS Accredited |  | Calibration Certificate Number: 13201 |  |
| Calibration Certificate Date: 15th March 2013 |  | Sample Stabilisation Time (minutes):60 |  |
| Reference Photometer Calibration Uncertainty: ± 2.4% (k=2, 20-200 lux, CIE illuminant A source) |  |  |  |
| Scan Set Up |  |  |  |
| Direction | Range |  | Increment |
| Inclination Zone 1 | 0-180° |  | 3° |
| Azimuth | 0-360° |  | 10° |
| Results |  |  |  |
| Integrated Luminous Flux (lumens):657.7 | Peak Intensity (3° Spot, candelas): 126.0 |  | Efficacy (lumens/Watt): 61.877777 |
| Beam Angle (50% of max intensity C0-180, degrees): 110.2 |  |  |  |
| Photometric Filename (IES LM-63-2002): 7134 MASHIKO LED 600 |  |  |  |
| IES File – Absolute or Relative Format? Absolute |  |  |  |
| Photometric Filename (EULUMDAT): 7134 MASHIKO LED 600 |  |  |  |
| EULUMDAT File – Absolute or Relative Format? Absolute |  |  |  |

Figure 3: Far-Field Luminous Intensity (C0-180, Cartesian Coordinates)

Figure 4: Far-Field Luminous Intensity (C0-180, C90-270, Polar Coordinates)

Figure 4. Illuminance cone diagram.

Table 1. UGR values

|  | 0 | 10 | 20 | 30 | 40 | 50 | 60 | 70 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150 | 160 | 170 | 180 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 |
| 5 | 118 | 120 | 121 | 123 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 124 | 124 | 123 | 122 | 120 | 119 |
| 10 | 117 | 120 | 122 | 121 | 122 | 123 | 124 | 125 | 126 | 126 | 125 | 125 | 124 | 123 | 122 | 122 | 121 | 120 | 119 |
| 15 | 114 | 118 | 119 | 120 | 122 | 124 | 124 | 124 | 125 | 125 | 125 | 125 | 125 | 125 | 122 | 120 | 119 | 119 | 115 |
| 20 | 111 | 116 | 117 | 117 | 118 | 121 | 121 | 123 | 124 | 124 | 124 | 124 | 124 | 122 | 121 | 119 | 115 | 117 | 111 |
| 25 | 105 | 110 | 112 | 112 | 117 | 118 | 120 | 122 | 122 | 122 | 122 | 123 | 120 | 117 | 116 | 115 | 112 | 110 | 104 |
| 30 | 99 | 103 | 107 | 110 | 112 | 112 | 117 | 120 | 121 | 123 | 122 | 119 | 117 | 116 | 110 | 109 | 106 | 104 | 97 |
| 35 | 92 | 96 | 100 | 104 | 105 | 110 | 112 | 117 | 119 | 122 | 119 | 117 | 116 | 110 | 106 | 102 | 100 | 96 | 91 |
| 40 | 84 | 88 | 92 | 97 | 99 | 104 | 111 | 115 | 116 | 116 | 115 | 114 | 111 | 109 | 102 | 96 | 93 | 90 | 84 |
| 45 | 76 | 80 | 85 | 90 | 95 | 100 | 107 | 111 | 112 | 115 | 114 | 111 | 107 | 103 | 96 | 89 | 86 | 82 | 76 |
| 50 | 68 | 71 | 78 | 83 | 89 | 94 | 102 | 107 | 108 | 110 | 105 | 103 | 101 | 97 | 90 | 82 | 78 | 73 | 67 |
| 55 | 57 | 62 | 70 | 75 | 82 | 89 | 96 | 100 | 103 | 103 | 103 | 101 | 94 | 88 | 83 | 75 | 69 | 64 | 57 |
| 60 | 47 | 52 | 62 | 68 | 75 | 80 | 91 | 92 | 96 | 98 | 96 | 92 | 90 | 81 | 75 | 68 | 61 | 54 | 47 |
| 65 | 37 | 43 | 53 | 61 | 68 | 73 | 83 | 87 | 91 | 91 | 90 | 87 | 82 | 76 | 67 | 61 | 53 | 44 | 38 |
| 70 | 27 | 35 | 44 | 53 | 60 | 66 | 75 | 79 | 83 | 84 | 83 | 79 | 75 | 68 | 60 | 53 | 43 | 34 | 29 |
| 75 | 18 | 25 | 35 | 45 | 52 | 60 | 67 | 72 | 75 | 77 | 74 | 72 | 66 | 59 | 53 | 44 | 34 | 25 | 19 |
| 80 | 8 | 15 | 27 | 37 | 46 | 53 | 58 | 64 | 68 | 68 | 67 | 62 | 58 | 52 | 44 | 35 | 25 | 15 | 9 |
| 85 | 5 | 11 | 22 | 32 | 40 | 47 | 54 | 57 | 61 | 61 | 60 | 57 | 52 | 45 | 38 | 29 | 20 | 11 | 5 |
| 90 | 1 | 6 | 18 | 27 | 35 | 44 | 49 | 54 | 55 | 57 | 56 | 53 | 47 | 40 | 33 | 24 | 15 | 6 | 1 |
| 95 | 1 | 5 | 16 | 25 | 33 | 38 | 44 | 51 | 48 | 50 | 54 | 51 | 46 | 39 | 32 | 23 | 14 | 6 | 1 |
| 100 | 1 | 5 | 14 | 22 | 30 | 37 | 45 | 49 | 52 | 53 | 53 | 50 | 45 | 39 | 31 | 22 | 14 | 5 | 1 |
| 105 | 1 | 5 | 13 | 21 | 28 | 36 | 44 | 47 | 51 | 51 | 51 | 48 | 43 | 39 | 30 | 22 | 14 | 5 | 1 |
| 110 | 1 | 5 | 12 | 20 | 27 | 35 | 41 | 46 | 48 | 49 | 49 | 45 | 41 | 36 | 29 | 22 | 13 | 5 | 1 |
| 115 | 1 | 5 | 12 | 20 | 26 | 33 | 39 | 43 | 46 | 48 | 46 | 43 | 38 | 34 | 27 | 20 | 12 | 5 | 1 |
| 120 | 1 | 5 | 11 | 18 | 25 | 33 | 36 | 40 | 43 | 44 | 43 | 41 | 36 | 31 | 26 | 19 | 12 | 5 | 1 |
| 125 | 1 | 5 | 10 | 17 | 23 | 29 | 35 | 37 | 40 | 41 | 39 | 38 | 35 | 30 | 24 | 17 | 11 | 4 | 1 |
| 130 | 0 | 4 | 9 | 16 | 22 | 28 | 31 | 35 | 37 | 37 | 37 | 35 | 31 | 27 | 22 | 16 | 11 | 4 | 1 |
| 135 | 1 | 4 | 9 | 14 | 19 | 25 | 29 | 32 | 33 | 34 | 33 | 32 | 29 | 25 | 20 | 15 | 9 | 4 | 1 |
| 140 | 2 | 3 | 8 | 13 | 18 | 22 | 26 | 28 | 30 | 31 | 30 | 28 | 26 | 22 | 18 | 14 | 8 | 4 | 1 |
| 145 | 1 | 3 | 7 | 11 | 15 | 19 | 23 | 25 | 26 | 26 | 26 | 25 | 22 | 20 | 16 | 12 | 7 | 3 | 1 |
| 150 | 1 | 3 | 6 | 10 | 13 | 16 | 19 | 21 | 23 | 24 | 23 | 21 | 19 | 17 | 14 | 10 | 6 | 3 | 2 |
| 155 | 1 | 3 | 5 | 8 | 11 | 14 | 16 | 18 | 19 | 19 | 18 | 17 | 16 | 14 | 11 | 7 | 4 | 3 | 1 |
| 160 | 1 | 2 | 4 | 6 | 9 | 11 | 13 | 14 | 14 | 15 | 14 | 13 | 12 | 10 | 8 | 6 | 4 | 2 | 1 |
| 165 | 2 | 2 | 3 | 5 | 6 | 7 | 9 | 10 | 10 | 10 | 10 | 10 | 9 | 7 | 6 | 5 | 3 | 2 | 1 |
| 170 | 2 | 2 | 3 | 4 | 4 | 5 | 5 | 6 | 6 | 7 | 6 | 6 | 6 | 5 | 4 | 4 | 3 | 2 | 1 |
| 175 | 2 | 2 | 2 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2 | 2 | 2 | 2 |
| 180 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |

Table 2a. Luminous intensity values, azimuth 0-180°

|  | 190 | 200 | 210 | 220 | 230 | 240 | 250 | 260 | 270 | 280 | 290 | 300 | 310 | 320 | 330 | 340 | 350 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 | 118 |
| 5 | 120 | 122 | 123 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 125 | 124 | 124 | 123 | 121 | 119 |
| 10 | 120 | 122 | 121 | 122 | 123 | 124 | 125 | 126 | 126 | 125 | 125 | 124 | 123 | 122 | 122 | 121 | 119 |
| 15 | 118 | 119 | 120 | 122 | 124 | 124 | 124 | 125 | 125 | 125 | 125 | 125 | 125 | 122 | 120 | 119 | 118 |
| 20 | 116 | 117 | 117 | 118 | 121 | 121 | 123 | 124 | 124 | 124 | 124 | 124 | 122 | 121 | 119 | 115 | 117 |
| 25 | 110 | 112 | 112 | 117 | 118 | 120 | 122 | 122 | 122 | 122 | 123 | 120 | 117 | 116 | 115 | 112 | 110 |
| 30 | 103 | 107 | 110 | 112 | 112 | 117 | 120 | 121 | 123 | 122 | 119 | 117 | 116 | 110 | 109 | 106 | 104 |
| 35 | 96 | 100 | 104 | 105 | 110 | 112 | 117 | 119 | 122 | 119 | 117 | 116 | 110 | 106 | 102 | 100 | 96 |
| 40 | 88 | 92 | 97 | 99 | 104 | 111 | 115 | 116 | 116 | 115 | 114 | 111 | 109 | 102 | 96 | 93 | 90 |
| 45 | 80 | 85 | 90 | 95 | 100 | 107 | 111 | 112 | 115 | 114 | 111 | 107 | 103 | 96 | 89 | 86 | 82 |
| 50 | 71 | 78 | 83 | 89 | 94 | 102 | 107 | 108 | 110 | 105 | 103 | 101 | 97 | 90 | 82 | 78 | 73 |
| 55 | 62 | 70 | 75 | 82 | 89 | 96 | 100 | 103 | 103 | 103 | 101 | 94 | 88 | 83 | 75 | 69 | 64 |
| 60 | 52 | 62 | 68 | 75 | 80 | 91 | 92 | 96 | 98 | 96 | 92 | 90 | 81 | 75 | 68 | 61 | 54 |
| 65 | 43 | 53 | 61 | 68 | 73 | 83 | 87 | 91 | 91 | 90 | 87 | 82 | 76 | 67 | 61 | 53 | 44 |
| 70 | 35 | 44 | 53 | 60 | 66 | 75 | 79 | 83 | 84 | 83 | 79 | 75 | 68 | 60 | 53 | 43 | 34 |
| 75 | 25 | 35 | 45 | 52 | 60 | 67 | 72 | 75 | 77 | 74 | 72 | 66 | 59 | 53 | 44 | 34 | 25 |
| 80 | 15 | 27 | 37 | 46 | 53 | 58 | 64 | 68 | 68 | 67 | 62 | 58 | 52 | 44 | 35 | 25 | 15 |
| 85 | 11 | 22 | 32 | 40 | 47 | 54 | 57 | 61 | 61 | 60 | 57 | 52 | 45 | 38 | 29 | 20 | 11 |
| 90 | 6 | 18 | 27 | 35 | 44 | 49 | 54 | 55 | 57 | 56 | 53 | 47 | 40 | 33 | 24 | 15 | 6 |
| 95 | 5 | 16 | 25 | 33 | 38 | 44 | 51 | 48 | 50 | 54 | 51 | 46 | 39 | 32 | 23 | 14 | 6 |
| 100 | 5 | 14 | 22 | 30 | 37 | 45 | 49 | 52 | 53 | 53 | 50 | 45 | 39 | 31 | 22 | 14 | 5 |
| 105 | 5 | 13 | 21 | 28 | 36 | 44 | 47 | 51 | 51 | 51 | 48 | 43 | 39 | 30 | 22 | 14 | 5 |
| 110 | 5 | 12 | 20 | 27 | 35 | 41 | 46 | 48 | 49 | 49 | 45 | 41 | 36 | 29 | 22 | 13 | 5 |
| 115 | 5 | 12 | 20 | 26 | 33 | 39 | 43 | 46 | 48 | 46 | 43 | 38 | 34 | 27 | 21 | 12 | 5 |
| 120 | 5 | 11 | 18 | 25 | 33 | 36 | 40 | 43 | 44 | 43 | 41 | 36 | 31 | 26 | 19 | 12 | 5 |
| 125 | 5 | 10 | 17 | 23 | 29 | 35 | 37 | 40 | 41 | 39 | 38 | 35 | 30 | 24 | 17 | 11 | 4 |
| 130 | 4 | 9 | 16 | 22 | 28 | 31 | 35 | 37 | 37 | 37 | 35 | 31 | 27 | 22 | 16 | 11 | 4 |
| 135 | 4 | 9 | 14 | 19 | 25 | 29 | 32 | 33 | 34 | 33 | 32 | 29 | 25 | 20 | 15 | 9 | 4 |
| 140 | 3 | 8 | 13 | 18 | 22 | 26 | 28 | 30 | 31 | 30 | 28 | 26 | 22 | 18 | 14 | 8 | 4 |
| 145 | 3 | 7 | 11 | 15 | 19 | 23 | 25 | 26 | 26 | 26 | 25 | 22 | 20 | 16 | 12 | 7 | 3 |
| 150 | 3 | 6 | 10 | 13 | 16 | 19 | 21 | 23 | 24 | 23 | 21 | 19 | 17 | 14 | 10 | 6 | 3 |
| 155 | 3 | 5 | 8 | 11 | 14 | 16 | 18 | 19 | 19 | 18 | 17 | 16 | 14 | 11 | 7 | 4 | 3 |
| 160 | 2 | 4 | 6 | 9 | 11 | 13 | 14 | 14 | 15 | 14 | 13 | 12 | 10 | 8 | 6 | 4 | 2 |
| 165 | 2 | 3 | 5 | 6 | 7 | 9 | 10 | 10 | 10 | 10 | 10 | 9 | 7 | 6 | 5 | 3 | 2 |
| 170 | 2 | 3 | 4 | 4 | 5 | 5 | 6 | 6 | 7 | 6 | 6 | 6 | 5 | 4 | 4 | 3 | 2 |
| 175 | 2 | 2 | 2 | 2 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2 | 2 | 2 |
| 180 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |

Table 2b. Luminous intensity values, azimuth 190-350°

Signature:

_________________________________

Print Name:

GH JOHN

_________________________________

Date:

_________________________________

Partner / Director
Duly authorised to sign on behalf of:

Photometric and Optical Testing Services LLP